GMP CSV Controlled Document

System Classification and Criticality Assessment

QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project

| Control Field | Value |
| --- | --- |
| Document number | QCL-2026-CLA-001 |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Risk and Compliance |
| Version | 1.0 |
| Effective/record date | 2026-03-20 |
| Status | Approved in fictional case |

Use notice: This Markdown file is a fully populated fictional CSV training case. It is not an executed GMP record and does not replace company procedures, approved signatures, supplier evidence, or site-specific risk decisions.

# Document control

| Field | Value |
| --- | --- |
| Document number | QCL-2026-CLA-001 |
| Company | NovaSterile Pharma Co., Ltd. |
| Site | Suzhou Sterile Products Manufacturing Site |
| Project | QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Risk and Compliance |
| Version | 1.0 |
| Effective/record date | 2026-03-20 |
| Case status | Approved in fictional case |

# Governing references

China GMP (2010 Revision), Appendix: Computerised Systems (effective 1 December 2015)

China GMP (2010 Revision), Appendix: Qualification and Validation (effective 1 December 2015)

EU GMP Annex 11: Computerised Systems (2011)

EU GMP Annex 15: Qualification and Validation (2015)

21 CFR Part 11: Electronic Records; Electronic Signatures

FDA Guidance: Part 11, Electronic Records; Electronic Signatures — Scope and Application

PIC/S PI 041-1: Good Practices for Data Management and Integrity in Regulated GMP/GDP Environments (2021)

ISPE GAMP 5, Second Edition (2022), used as non-binding industry guidance

# 1. Classification method

Components are classified by their role and complexity. Industry categories are used as a planning aid, not as a substitute for GxP risk assessment.

# 2. Classification results

| Component | Classification | Configuration/customisation | Criticality | Validation implication |
| --- | --- | --- | --- | --- |
| LabSphere LIMS/LES 8.4.2 | Configurable commercial application | Extensive workflow, master data, reports and security configuration | High | URS, configuration specification, risk-based OQ/PQ and traceability |
| QCL-IGW-01 OS/runtime | Infrastructure software | Standard hardened configuration | High supporting | Installation qualification and operational monitoring |
| SAP/eQMS/CDS connectors | Custom software/integration | Custom mappings, transformations and retry logic | High | Development plan, code review, unit and interface testing |
| Report templates/calculation rules | Configured/custom content | Site-specific logic | High | Specification, independent verification and regression |
| Barcode printers/scanners | Hardware/peripheral | Standard drivers and device mapping | Medium | Installation and workflow test |
| ArchiveVault | Configurable archive service | Retention/index configuration | High | Supplier assessment, qualification and retrieval test |
| Controlled spreadsheets | User-developed applications | Formula/macro dependent by file | Medium/High | Individual assessment, migration or retirement evidence |

# 3. Criticality evaluation

The application is business-critical and data-integrity-critical. An incorrect result, limit, signature or interface transfer could support an incorrect disposition. Loss of availability beyond eight hours delays QC release. Loss of historical records impairs inspection and product-quality investigations.

# 4. Validation strategy decision

Use a full risk-based lifecycle. Supplier standard tests may be leveraged for non-site-specific product functions. All site configuration, interfaces, calculations, roles, audit trails, signatures, migration and representative business processes require company-controlled verification.

# 5. Related documents

| Relationship | Document ID | Document |
| --- | --- | --- |
| Input | QCL-2026-IUS-001 | [QCLabOne 2.0 Intended Use Statement](007_Intended_Use_Statement.md) |
| Input | QCL-2026-GXP-001 | [QCLabOne 2.0 GxP Impact Assessment](009_GxP_Impact_Assessment.md) |
| Output | QCL-2026-PRA-001 | [Preliminary Risk Assessment](011_Preliminary_Risk_Assessment.md) |
| Output | QCL-2026-VP-001 | [QCLabOne 2.0 Project Validation Plan](015_Project_Validation_Plan.md) |
| Output | QCL-2026-SDL-001 | [Supplier Documentation Leveraging Assessment](021_Supplier_Documentation_Leveraging_Assessment.md) |
| Output | QCL-2026-FRA-001 | [Functional Risk Assessment](038_Functional_Risk_Assessment.md) |
| Output | QCL-2026-VTP-001 | [Validation Test Plan and Strategy](046_Validation_Test_Plan_and_Strategy.md) |

# Approval record

| Approval step | Role | Case outcome |
| --- | --- | --- |
| Prepared by | CSV/Quality Risk Lead | Completed |
| Reviewed by | System Owner | Completed |
| Approved by | QA Validation Manager | Approved in fictional case |

# Revision history

| Version | Date | Change |
| --- | --- | --- |
| 1.0 | 2026-03-20 | Initial approved fictional case version. |