GMP CSV Controlled Document

Site Acceptance Test Protocol and Report

QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project

| Control Field | Value |
| --- | --- |
| Document number | QCL-2026-SAT-001 |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Verification |
| Version | 1.0 |
| Effective/record date | 2027-02-20 |
| Status | Approved in fictional case |

Use notice: This Markdown file is a fully populated fictional CSV training case. It is not an executed GMP record and does not replace company procedures, approved signatures, supplier evidence, or site-specific risk decisions.

# Document control

| Field | Value |
| --- | --- |
| Document number | QCL-2026-SAT-001 |
| Company | NovaSterile Pharma Co., Ltd. |
| Site | Suzhou Sterile Products Manufacturing Site |
| Project | QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Verification |
| Version | 1.0 |
| Effective/record date | 2027-02-20 |
| Case status | Approved in fictional case |

# Governing references

China GMP (2010 Revision), Appendix: Computerised Systems (effective 1 December 2015)

China GMP (2010 Revision), Appendix: Qualification and Validation (effective 1 December 2015)

EU GMP Annex 11: Computerised Systems (2011)

EU GMP Annex 15: Qualification and Validation (2015)

21 CFR Part 11: Electronic Records; Electronic Signatures

FDA Guidance: Part 11, Electronic Records; Electronic Signatures — Scope and Application

PIC/S PI 041-1: Good Practices for Data Management and Integrity in Regulated GMP/GDP Environments (2021)

ISPE GAMP 5, Second Edition (2022), used as non-binding industry guidance

# 1. Objective

SAT confirmed that the approved solution operated correctly with site networks, identity, interfaces, instruments, peripherals, monitoring and archive services.

# 2. Execution

Executed 2027-02-13 to 2027-02-18 in the site validation environment using the approved deployment package.

# 3. Results

| Test | Area | Acceptance | Result |
| --- | --- | --- | --- |
| SAT-001 | Production/validation URL and network path | Accessible only from approved networks | Pass |
| SAT-002 | Corporate SSO and role claims | Correct identity/role mapping | Pass |
| SAT-003 | SAP connectivity and sample message | End-to-end acknowledgement | Pass |
| SAT-004 | eQMS/CDS connectivity | Secure connection and sandbox transaction | Pass |
| SAT-005 | 24 instrument/device endpoints | Correct device identity and normalized data | Pass |
| SAT-006 | Barcode printers/scanners | Readable, correctly associated barcode | Pass after DEF-006 |
| SAT-007 | Monitoring, logs and alerts | Alerts reach approved support groups | Pass |
| SAT-008 | Backup and gateway configuration export | Successful and monitored | Pass |
| SAT-009 | Vendor remote support workflow | Disabled by default; approved session logged | Pass |
| SAT-010 | ArchiveVault test package | Transferred, checksum and retrieved | Pass |

# 4. Deviation/defect

One printer had an obsolete driver (DEF-006). The approved driver was installed and the affected print/scan test repeated successfully.

# 5. Conclusion

The site environment met SAT exit criteria and was released to formal qualification/validation testing.

# 6. Related documents

| Relationship | Document ID | Document |
| --- | --- | --- |
| Input | QCL-2026-BLD-001 | [Installation, Build and Configuration Record](040_Installation_Build_and_Configuration_Record.md) |
| Input | QCL-2026-CIR-001 | [Configuration Item and Version Register](041_Configuration_Item_and_Version_Register.md) |
| Input | QCL-2026-FAT-001 | [Factory Acceptance Test Protocol and Report](044_Factory_Acceptance_Test_Protocol_and_Report.md) |
| Output | QCL-2026-VTP-001 | [Validation Test Plan and Strategy](046_Validation_Test_Plan_and_Strategy.md) |
| Output | QCL-2026-IQ-001 | [Installation Qualification Protocol and Report](048_Installation_Qualification_Protocol_and_Report.md) |
| Output | QCL-2026-IFT-001 | [Interface Test Protocol and Report](051_Interface_Test_Protocol_and_Report.md) |
| Output | QCL-2026-DEF-001 | [Defect Log and Resolution Report](060_Defect_Log_and_Resolution_Report.md) |

# Approval record

| Approval step | Role | Case outcome |
| --- | --- | --- |
| Prepared by | Test Lead | Completed |
| Reviewed by | Business Process Owner | Completed |
| Approved by | QA Validation Manager | Approved in fictional case |

# Revision history

| Version | Date | Change |
| --- | --- | --- |
| 1.0 | 2027-02-20 | Initial approved fictional case version. |