GMP CSV Controlled Document

Audit Trail Challenge Test Report

QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project

| Control Field | Value |
| --- | --- |
| Document number | QCL-2026-ATC-001 |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Verification |
| Version | 1.0 |
| Effective/record date | 2027-04-30 |
| Status | Approved in fictional case |

Use notice: This Markdown file is a fully populated fictional CSV training case. It is not an executed GMP record and does not replace company procedures, approved signatures, supplier evidence, or site-specific risk decisions.

# Document control

| Field | Value |
| --- | --- |
| Document number | QCL-2026-ATC-001 |
| Company | NovaSterile Pharma Co., Ltd. |
| Site | Suzhou Sterile Products Manufacturing Site |
| Project | QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Verification |
| Version | 1.0 |
| Effective/record date | 2027-04-30 |
| Case status | Approved in fictional case |

# Governing references

China GMP (2010 Revision), Appendix: Computerised Systems (effective 1 December 2015)

China GMP (2010 Revision), Appendix: Qualification and Validation (effective 1 December 2015)

EU GMP Annex 11: Computerised Systems (2011)

EU GMP Annex 15: Qualification and Validation (2015)

21 CFR Part 11: Electronic Records; Electronic Signatures

FDA Guidance: Part 11, Electronic Records; Electronic Signatures — Scope and Application

PIC/S PI 041-1: Good Practices for Data Management and Integrity in Regulated GMP/GDP Environments (2021)

ISPE GAMP 5, Second Edition (2022), used as non-binding industry guidance

# 1. Objective

Challenge audit-trail generation, completeness, protection, reason-for-change, review and export for critical GxP objects.

# 2. Results

| Test | Scenario | Actual result | Outcome |
| --- | --- | --- | --- |
| ATC-001 | Result creation/change/delete attempt | Create and change recorded with old/new value; delete denied and attempt logged | Pass |
| ATC-002 | Specification/formula/master-data changes | All critical fields recorded | Pass after DEV-001 |
| ATC-003 | Mandatory reason for critical change | Save blocked without reason | Pass after DEV-001 |
| ATC-004 | Protection/filter/export/review | No tested role could alter trail; export matched source/filter | Pass |
| ATC-005 | Role/signature/workflow events | Actor, target, old/new state, time and meaning captured | Pass |

# 3. Deviation DEV-001

Initial configuration did not require a reason when changing a sample comment classified by the site as critical metadata. Event classification and mandatory-reason configuration were corrected in package `QCL-CONFIG-1.0-R1`. ATC-002, ATC-003 and affected regression tests passed.

# 4. Review usability

Reviewers successfully filtered by record, user, date, event type and changed field; exported evidence identified filters and execution time. No ordinary or routine administrative role could edit/delete audit entries.

# 5. Conclusion

Audit-trail controls are acceptable for intended use and operational review.

# 6. Related documents

| Relationship | Document ID | Document |
| --- | --- | --- |
| Input | QCL-2026-DIRA-001 | [Data Integrity Risk Assessment](012_Data_Integrity_Risk_Assessment.md) |
| Input | QCL-2026-ERES-001 | [21 CFR Part 11 and EU Annex 11 Assessment](013_21_CFR_Part_11_and_EU_Annex_11_Assessment.md) |
| Input | QCL-2026-CS-001 | [Configuration Specification](025_Configuration_Specification.md) |
| Input | QCL-2026-ATS-001 | [Audit Trail Specification](032_Audit_Trail_Specification.md) |
| Input | QCL-2026-FRA-001 | [Functional Risk Assessment](038_Functional_Risk_Assessment.md) |
| Input | QCL-2026-VTP-001 | [Validation Test Plan and Strategy](046_Validation_Test_Plan_and_Strategy.md) |
| Input | QCL-2026-MTS-001 | [Master Test Scripts and Test Data Catalogue](047_Master_Test_Scripts_and_Test_Data_Catalogue.md) |
| Input | QCL-2026-OQ-001 | [Operational Qualification Protocol and Report](049_Operational_Qualification_Protocol_and_Report.md) |
| Output | QCL-2026-DEV-001 | [Validation Deviation Log and Investigation Report](061_Validation_Deviation_Log_and_Investigation_Report.md) |
| Output | QCL-2026-CAPA-001 | [CAPA Record and Effectiveness Check](062_CAPA_Record_and_Effectiveness_Check.md) |
| Output | QCL-2026-TSR-001 | [Test Summary Report](063_Test_Summary_Report.md) |
| Output | QCL-2026-RTM-002 | [Final Requirements Traceability Matrix](064_Final_Requirements_Traceability_Matrix.md) |
| Output | QCL-2026-VSR-001 | [Validation Summary Report](065_Validation_Summary_Report.md) |
| Output | SOP-QA-ATR-001 | [QCLabOne 2.0 Audit Trail Review SOP](068_Audit_Trail_Review_SOP.md) |
| Output | QCL-2027-ATR-001 | [Periodic Audit Trail Review Record](081_Periodic_Audit_Trail_Review_Record.md) |
| Output | QCL-2027-DIM-001 | [Data Integrity Monitoring and Trend Report](082_Data_Integrity_Monitoring_and_Trend_Report.md) |

# Approval record

| Approval step | Role | Case outcome |
| --- | --- | --- |
| Prepared by | Test Lead | Completed |
| Reviewed by | Business Process Owner | Completed |
| Approved by | QA Validation Manager | Approved in fictional case |

# Revision history

| Version | Date | Change |
| --- | --- | --- |
| 1.0 | 2027-04-30 | Initial approved fictional case version. |