GMP CSV Controlled Document

Performance, Load and Capacity Test Report

QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project

| Control Field | Value |
| --- | --- |
| Document number | QCL-2026-PLC-001 |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Verification |
| Version | 1.0 |
| Effective/record date | 2027-04-30 |
| Status | Approved in fictional case |

Use notice: This Markdown file is a fully populated fictional CSV training case. It is not an executed GMP record and does not replace company procedures, approved signatures, supplier evidence, or site-specific risk decisions.

# Document control

| Field | Value |
| --- | --- |
| Document number | QCL-2026-PLC-001 |
| Company | NovaSterile Pharma Co., Ltd. |
| Site | Suzhou Sterile Products Manufacturing Site |
| Project | QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Verification |
| Version | 1.0 |
| Effective/record date | 2027-04-30 |
| Case status | Approved in fictional case |

# Governing references

China GMP (2010 Revision), Appendix: Computerised Systems (effective 1 December 2015)

China GMP (2010 Revision), Appendix: Qualification and Validation (effective 1 December 2015)

EU GMP Annex 11: Computerised Systems (2011)

EU GMP Annex 15: Qualification and Validation (2015)

21 CFR Part 11: Electronic Records; Electronic Signatures

FDA Guidance: Part 11, Electronic Records; Electronic Signatures — Scope and Application

PIC/S PI 041-1: Good Practices for Data Management and Integrity in Regulated GMP/GDP Environments (2021)

ISPE GAMP 5, Second Edition (2022), used as non-binding industry guidance

# 1. Objective

Confirm that the production design supports expected users, interface throughput, reports and five-year data growth.

# 2. Results

| Test | Load/scenario | Acceptance | Actual/outcome |
| --- | --- | --- | --- |
| PLC-001 | 200 concurrent mixed users | 95th percentile ≤3.0 s for standard transactions | 2.4 s after cache/index correction; Pass |
| PLC-002 | 10,000 interface messages/hour | No loss; backlog clears within 30 min | 10,000 accepted/rejected as designed; backlog 12 min; Pass |
| PLC-003 | 100,000-result report | Generate ≤5 min without truncation | 3 min 42 sec; Pass |
| PLC-004 | Five-year capacity forecast | ≥30% headroom and alerts at 70/85% | Projected peak 61%; alert tests passed |

# 3. Deviation DEV-004

The initial stability dashboard response was 3.8 seconds at peak load. A database index and cache setting reduced the 95th percentile to 2.4 seconds. No data correctness or integrity issue occurred.

# 4. Conclusion

Performance and capacity meet the approved release criteria.

# 5. Related documents

| Relationship | Document ID | Document |
| --- | --- | --- |
| Input | QCL-2026-IAD-001 | [Infrastructure and Architecture Design](026_Infrastructure_and_Architecture_Design.md) |
| Input | QCL-2026-RCS-001 | [Report and Calculation Specification](029_Report_and_Calculation_Specification.md) |
| Input | QCL-2026-FRA-001 | [Functional Risk Assessment](038_Functional_Risk_Assessment.md) |
| Input | QCL-2026-VTP-001 | [Validation Test Plan and Strategy](046_Validation_Test_Plan_and_Strategy.md) |
| Input | QCL-2026-MTS-001 | [Master Test Scripts and Test Data Catalogue](047_Master_Test_Scripts_and_Test_Data_Catalogue.md) |
| Output | QCL-2026-DEV-001 | [Validation Deviation Log and Investigation Report](061_Validation_Deviation_Log_and_Investigation_Report.md) |
| Output | QCL-2026-TSR-001 | [Test Summary Report](063_Test_Summary_Report.md) |
| Output | QCL-2026-RTM-002 | [Final Requirements Traceability Matrix](064_Final_Requirements_Traceability_Matrix.md) |
| Output | QCL-2026-VSR-001 | [Validation Summary Report](065_Validation_Summary_Report.md) |
| Output | QCL-2028-PRR-001 | [Annual Periodic Review Report](092_Annual_Periodic_Review_Report.md) |

# Approval record

| Approval step | Role | Case outcome |
| --- | --- | --- |
| Prepared by | Test Lead | Completed |
| Reviewed by | Business Process Owner | Completed |
| Approved by | QA Validation Manager | Approved in fictional case |

# Revision history

| Version | Date | Change |
| --- | --- | --- |
| 1.0 | 2027-04-30 | Initial approved fictional case version. |