GMP CSV Controlled Document

Data Retention, Archiving and Record Retrieval SOP

QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project

| Control Field | Value |
| --- | --- |
| Document number | SOP-QA-ARC-001 |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | SOP and Training |
| Version | 1.0 |
| Effective/record date | 2027-05-31 |
| Status | Approved in fictional case |

Use notice: This Markdown file is a fully populated fictional CSV training case. It is not an executed GMP record and does not replace company procedures, approved signatures, supplier evidence, or site-specific risk decisions.

# Document control

| Field | Value |
| --- | --- |
| Document number | SOP-QA-ARC-001 |
| Company | NovaSterile Pharma Co., Ltd. |
| Site | Suzhou Sterile Products Manufacturing Site |
| Project | QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | SOP and Training |
| Version | 1.0 |
| Effective/record date | 2027-05-31 |
| Case status | Approved in fictional case |

# Governing references

China GMP (2010 Revision), Appendix: Computerised Systems (effective 1 December 2015)

China GMP (2010 Revision), Appendix: Qualification and Validation (effective 1 December 2015)

EU GMP Annex 11: Computerised Systems (2011)

EU GMP Annex 15: Qualification and Validation (2015)

21 CFR Part 11: Electronic Records; Electronic Signatures

FDA Guidance: Part 11, Electronic Records; Electronic Signatures — Scope and Application

PIC/S PI 041-1: Good Practices for Data Management and Integrity in Regulated GMP/GDP Environments (2021)

ISPE GAMP 5, Second Edition (2022), used as non-binding industry guidance

# 1. Purpose

This SOP controls retention, archive transfer, integrity verification, authorised retrieval and final disposition of QCLabOne and Legacy LIMS records.

# 2. Retention schedule

| Record class | Retention rule | Repository | Owner |
| --- | --- | --- | --- |
| Approved QC sample/test/result record | Product/batch retention requirement; at least 1 year after expiry unless longer rule applies | QCLabOne then ArchiveVault | QC Records Owner |
| Audit trail and electronic signature | Same as associated record | With record package | QA Data Integrity |
| Specification/method/calculation version | Effective life plus related record retention | QCLabOne/eDMS archive | QC Master Data Owner |
| OOS/OOT link and local summary | Same as related test/investigation | QCLabOne + eQMS | QA |
| Interface messages/reconciliation | 2 years online; longer when linked to record/incident | Log/archive repository | System Owner |
| Access/change/incident/backup evidence | System life + 2 years or longer if related to GMP investigation | eQMS/eDMS | System Owner/QA |
| Legacy LIMS archive | Original required retention by record class | ArchiveVault read-only | Records Management |
| Validation/retirement dossier index | Permanent index; evidence per applicable schedule | eDMS | QA Validation |

# 3. Archive package

A package contains PDF/A rendering where applicable, structured metadata/data, attachments, audit trail, signature manifestation, version references and SHA-256 manifest. Transfer is reconciled and the archive is read-only.

# 4. Retrieval

Authorised users search by sample ID, batch, material, date, test and legacy reference. Retrieval records requester, purpose and date. Inspection requests receive complete readable copies without altering originals.

# 5. Archive verification

Before source deletion or retirement, verify package counts/checksums, metadata relationships, sample retrieval, audit trail, attachments and signature context. Failures require incident/deviation and prevent decommissioning.

# 6. Disposition

No routine user can purge records. Final deletion after retention requires Records Management and Quality approval, legal-hold check, documented list, secure deletion evidence and inventory update.

# 7. Related documents

| Relationship | Document ID | Document |
| --- | --- | --- |
| Input | QCL-2026-DIRA-001 | [Data Integrity Risk Assessment](012_Data_Integrity_Risk_Assessment.md) |
| Input | QCL-2026-DDS-001 | [Data Flow, Data Dictionary and Metadata Specification](028_Data_Flow_Data_Dictionary_and_Metadata_Specification.md) |
| Input | QCL-2026-BRS-001 | [Backup and Restore Specification](034_Backup_and_Restore_Specification.md) |
| Input | QCL-2026-DMP-001 | [Data Migration Strategy and Plan](036_Data_Migration_Strategy_and_Plan.md) |
| Input | QCL-2026-VSR-001 | [Validation Summary Report](065_Validation_Summary_Report.md) |
| Output | QCL-2028-ARC-001 | [System and Data Archive Record](097_System_and_Data_Archive_Record.md) |
| Output | QCL-2028-IRD-001 | [Inspection Readiness Dossier and Audit Response Record](098_Inspection_Readiness_Dossier_and_Audit_Response_Record.md) |
| Output | QCL-2028-LAR-001 | [Legacy Data Archiving and Retention Plan](102_Legacy_Data_Archiving_and_Retention_Plan.md) |
| Output | QCL-2028-AVR-001 | [Legacy Archive Verification and Retrievability Report](103_Legacy_Archive_Verification_and_Retrievability_Report.md) |
| Output | QCL-2028-RSR-001 | [Legacy LIMS Retirement Summary Report](106_Legacy_LIMS_Retirement_Summary_Report.md) |
| Output | QCL-2028-FDI-001 | [Final Validation and Retirement Dossier Index](108_Final_Validation_and_Retirement_Dossier_Index.md) |

# Approval record

| Approval step | Role | Case outcome |
| --- | --- | --- |
| Prepared by | Training/Procedure Owner | Completed |
| Reviewed by | System Owner | Completed |
| Approved by | QA Department Head | Approved in fictional case |

# Revision history

| Version | Date | Change |
| --- | --- | --- |
| 1.0 | 2027-05-31 | Initial approved fictional case version. |