GMP CSV Controlled Document

Functional Specification

QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project

| Control Field | Value |
| --- | --- |
| Document number | QCL-2026-FS-001 |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Design |
| Version | 1.0 |
| Effective/record date | 2026-10-31 |
| Status | Approved in fictional case |

Use notice: This Markdown file is a fully populated fictional CSV training case. It is not an executed GMP record and does not replace company procedures, approved signatures, supplier evidence, or site-specific risk decisions.

# Document control

| Field | Value |
| --- | --- |
| Document number | QCL-2026-FS-001 |
| Company | NovaSterile Pharma Co., Ltd. |
| Site | Suzhou Sterile Products Manufacturing Site |
| Project | QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Design |
| Version | 1.0 |
| Effective/record date | 2026-10-31 |
| Case status | Approved in fictional case |

# Governing references

China GMP (2010 Revision), Appendix: Computerised Systems (effective 1 December 2015)

China GMP (2010 Revision), Appendix: Qualification and Validation (effective 1 December 2015)

EU GMP Annex 11: Computerised Systems (2011)

EU GMP Annex 15: Qualification and Validation (2015)

21 CFR Part 11: Electronic Records; Electronic Signatures

FDA Guidance: Part 11, Electronic Records; Electronic Signatures — Scope and Application

PIC/S PI 041-1: Good Practices for Data Management and Integrity in Regulated GMP/GDP Environments (2021)

ISPE GAMP 5, Second Edition (2022), used as non-binding industry guidance

# 1. Purpose

This specification describes the functional behaviour that will satisfy the approved URS. Detailed implementation parameters are controlled in the design and configuration specifications.

# 2. Functional requirements

| Function ID | Function | URS coverage | Functional behaviour |
| --- | --- | --- | --- |
| FS-GEN-001 | Environment identification and production baseline control | URS-GEN-003, URS-GEN-005 | Application banner displays environment/version; deployment requires approved release package. |
| FS-GEN-002 | Unique identifiers and synchronized timestamps | URS-GEN-002, URS-GEN-004 | Database sequences generate immutable IDs; all services synchronize to corporate NTP. |
| FS-GEN-003 | In-scope QC business process enablement | URS-GEN-001 | Configured modules and workflows support chemistry, microbiology, stability and environmental monitoring. |
| FS-SMP-001 | SAP sample request intake | URS-SMP-001, URS-INT-001 | Gateway validates schema, checks idempotency and creates a pending request. |
| FS-SMP-002 | Sample ID and barcode generation | URS-SMP-002 | LIMS assigns QCL-YYYY-NNNNNNN and Code 128 barcode. |
| FS-SMP-003 | Receipt and chain of custody | URS-SMP-003, URS-SMP-004, URS-SMP-007 | Receipt and custody events require user, timestamp, location, quantity and reason where applicable. |
| FS-SMP-004 | Test assignment and sample completion | URS-SMP-005, URS-SMP-006 | Effective specification creates required tests; closure is blocked while mandatory tasks/investigations remain open. |
| FS-SMP-005 | Stability scheduling | URS-SMP-008 | Protocol-derived pull dates, windows and alerts are generated nightly. |
| FS-SPEC-001 | Specification and method lifecycle | URS-SPEC-001, URS-SPEC-002, URS-SPEC-003 | Draft-review-approve-effective-obsolete states with effective dating and immutable history. |
| FS-SPEC-002 | Result type and limit configuration | URS-SPEC-004 | Numeric, text, enumerated and qualitative types support configured validation and limits. |
| FS-SPEC-003 | Production promotion approval | URS-SPEC-005, URS-GEN-005 | Approved change record and electronic approval are prerequisites for promotion. |
| FS-LES-001 | Controlled electronic worksheet execution | URS-LES-001, URS-LES-002 | Only approved versions are executable; mandatory steps enforce sequence and completion. |
| FS-LES-002 | Resource references | URS-LES-003 | Reagent, standard, instrument and lot are selected from effective master data. |
| FS-LES-003 | Direct instrument capture | URS-LES-004 | Gateway associates device message with active test and retains source metadata. |
| FS-LES-004 | Manual entry and verification | URS-LES-005 | Manual critical values require reason and independent verifier. |
| FS-LES-005 | Pause/abort/repeat and exception workflow | URS-LES-006, URS-LES-008 | Controlled states require reason and supervisor action; dashboards show exceptions. |
| FS-LES-006 | Second-person review segregation | URS-LES-007 | Workflow engine prevents submitter from final review/approval. |
| FS-CALC-001 | Formula execution and unit handling | URS-CALC-001 | Versioned calculation service executes approved formulas with dimensional checks. |
| FS-CALC-002 | Rounding and significant figures | URS-CALC-002 | Rounding occurs only at configured display/final step using half-up rules unless specified. |
| FS-CALC-003 | Formula history and limit evaluation | URS-CALC-003, URS-CALC-004 | Record stores formula version and evaluates effective limit set. |
| FS-CALC-004 | Calculation change control | URS-CALC-005 | Formula changes require approved configuration package and regression evidence. |
| FS-INT-001 | SAP bidirectional interface | URS-INT-001 | Inbound sample and outbound status/result messages use unique message IDs and acknowledgements. |
| FS-INT-002 | eQMS OOS/OOT interface | URS-INT-002 | Idempotent REST message creates one case and stores returned case ID/status. |
| FS-INT-003 | CDS/SDMS approved-result interface | URS-INT-003 | Interface receives approved numeric result, unit, sequence ID, method version and stable source link. |
| FS-INT-004 | Instrument exception handling | URS-INT-004, URS-INT-005 | Malformed/duplicate messages enter a controlled queue with error and reprocessing history. |
| FS-INT-005 | Manual contingency processing | URS-INT-006 | Approved outage mode labels manually entered data and requires independent reconciliation. |
| FS-SEC-001 | Corporate authentication and individual identity | URS-SEC-001 | SAML federation uses unique corporate identity; local interactive accounts are disabled except break-glass accounts. |
| FS-SEC-002 | Role-based authorization | URS-SEC-002, URS-SEC-003 | Permissions are assigned through approved roles; admin and QC approval privileges are mutually exclusive. |
| FS-SEC-003 | Session and account lifecycle controls | URS-SEC-004, URS-SEC-005 | 15-minute lock; approved request controls create/change/disable actions. |
| FS-SEC-004 | Vendor remote support | URS-SEC-006 | Just-in-time access requires company ticket, MFA, time window and session logging. |
| FS-SEC-005 | Security event logging | URS-SEC-007 | Failed logins, lockouts, role changes and privileged actions are reportable. |
| FS-DI-001 | Protected audit trail | URS-DI-001, URS-DI-003 | Database audit service is always on; no application role can alter entries. |
| FS-DI-002 | Reason for change | URS-DI-002 | Critical changes require selection from controlled reason list plus optional comment. |
| FS-DI-003 | Audit trail review/export | URS-DI-004 | Authorized filters and PDF/CSV exports preserve source identifiers and criteria. |
| FS-DI-004 | Retention, deletion and archive | URS-DI-005, URS-DI-006 | Purge disabled for routine roles; archive job preserves data, metadata, audit and attachments. |
| FS-DI-005 | Transfer/migration reconciliation | URS-DI-007 | Count, hash, critical-field and relationship controls generate reconciliation reports. |
| FS-DI-006 | Time control | URS-DI-008 | NTP sync alarms at >60 seconds drift; timezone configuration is change-controlled. |
| FS-ESIG-001 | Signature identity and re-authentication | URS-ESIG-001, URS-ESIG-003 | Signature uses individual identity and credential re-entry for submit/review/approve. |
| FS-ESIG-002 | Signature manifestation and binding | URS-ESIG-002, URS-ESIG-004 | Name, time and meaning display; post-signature critical change invalidates approval and re-routes workflow. |
| FS-REP-001 | Controlled reports | URS-REP-001, URS-REP-002 | Approved templates generate reports with record and template metadata. |
| FS-REP-002 | PDF/print/export fidelity | URS-REP-003, URS-REP-004 | Rendered output retains units, pagination, signatures and query criteria. |
| FS-BCP-001 | Backup and restore | URS-BCP-001, URS-BCP-003 | Encrypted backups cover database, configuration and attachments; restore procedure verifies consistency. |
| FS-BCP-002 | Disaster recovery and continuity | URS-BCP-002, URS-BCP-004 | Standby region supports RPO 15 min/RTO 8 h; controlled manual forms support outage. |
| FS-MIG-001 | Migration and archive | URS-MIG-001, URS-MIG-002 | ETL and archive paths use approved mapping, validation and read-only retention. |
| FS-MIG-002 | Migration exceptions and source protection | URS-MIG-003, URS-MIG-004 | Exception log requires disposition; source remains read-only and backed up until acceptance. |
| FS-OPS-001 | Operational evidence and periodic review | URS-OPS-001, URS-OPS-003 | Reports and schedules support access/audit/backup review and annual validated-state assessment. |
| FS-OPS-002 | Supplier notification and retirement | URS-OPS-002, URS-OPS-004 | Quality agreement routes notices; retirement workflow requires archive and decommissioning approvals. |

# 3. Cross-cutting functional rules

All GxP state changes are attributable to an individual or approved service account.

Historical effective versions remain linked to completed records.

Interface processing is idempotent and exceptions are visible.

Post-signature changes invalidate affected approval as configured.

Routine users cannot delete GxP records or audit trails.

Application errors display a user-safe message and create diagnostic logs without exposing secrets.

Cross-cutting exception behavior shall be consistent across modules. When a required prerequisite is missing, the system shall prevent completion, identify the missing condition and retain the attempted action where required for auditability. When a technical fault occurs, the system shall avoid partial regulated records and route recoverable items to an exception queue with owner, timestamp and disposition.

| Functional rule type | Verification focus |
| --- | --- |
| Workflow prerequisite | Negative test confirms blocked action and clear reason |
| Historical version use | Completed records retain the approved effective version |
| Exception queue | Failed transactions are visible, attributable and dispositioned |
| Signature invalidation | Post-signature change forces re-review and re-signature |
| Error handling | User message is safe and diagnostic log supports investigation |

# 4. Related documents

| Relationship | Document ID | Document |
| --- | --- | --- |
| Input | QCL-2026-URS-001 | [User Requirements Specification](022_User_Requirements_Specification.md) |
| Input | QCL-2026-DIRA-001 | [Data Integrity Risk Assessment](012_Data_Integrity_Risk_Assessment.md) |
| Input | QCL-2026-SDL-001 | [Supplier Documentation Leveraging Assessment](021_Supplier_Documentation_Leveraging_Assessment.md) |
| Output | QCL-2026-DS-001 | [System Design Specification](024_System_Design_Specification.md) |
| Output | QCL-2026-CS-001 | [Configuration Specification](025_Configuration_Specification.md) |
| Output | QCL-2026-ICS-001 | [Interface Control Specification](027_Interface_Control_Specification.md) |
| Output | QCL-2026-RCS-001 | [Report and Calculation Specification](029_Report_and_Calculation_Specification.md) |
| Output | QCL-2026-MDS-001 | [Master Data Specification](030_Master_Data_Specification.md) |
| Output | QCL-2026-RAC-001 | [Role and Access Control Matrix](031_Role_and_Access_Control_Matrix.md) |
| Output | QCL-2026-ATS-001 | [Audit Trail Specification](032_Audit_Trail_Specification.md) |
| Output | QCL-2026-ESS-001 | [Electronic Signature Specification](033_Electronic_Signature_Specification.md) |
| Output | QCL-2026-FRA-001 | [Functional Risk Assessment](038_Functional_Risk_Assessment.md) |
| Output | QCL-2026-RTM-001 | [Initial Requirements Traceability Matrix](039_Initial_Requirements_Traceability_Matrix.md) |

# Approval record

| Approval step | Role | Case outcome |
| --- | --- | --- |
| Prepared by | System Owner | Completed |
| Reviewed by | Vendor Solution Architect | Completed |
| Approved by | QA CSV Lead | Approved in fictional case |

# Revision history

| Version | Date | Change |
| --- | --- | --- |
| 1.0 | 2026-10-31 | Initial approved fictional case version. |