GMP CSV Controlled Document

Validation Summary Report

QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project

| Control Field | Value |
| --- | --- |
| Document number | QCL-2026-VSR-001 |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Release |
| Version | 1.0 |
| Effective/record date | 2027-06-05 |
| Status | Approved in fictional case |

Use notice: This Markdown file is a fully populated fictional CSV training case. It is not an executed GMP record and does not replace company procedures, approved signatures, supplier evidence, or site-specific risk decisions.

# Document control

| Field | Value |
| --- | --- |
| Document number | QCL-2026-VSR-001 |
| Company | NovaSterile Pharma Co., Ltd. |
| Site | Suzhou Sterile Products Manufacturing Site |
| Project | QCLabOne 2.0 QC Laboratory Digital Platform Replacement Project |
| System | QCLabOne 2.0 LIMS/LES |
| Lifecycle phase | Release |
| Version | 1.0 |
| Effective/record date | 2027-06-05 |
| Case status | Approved in fictional case |

# Governing references

China GMP (2010 Revision), Appendix: Computerised Systems (effective 1 December 2015)

China GMP (2010 Revision), Appendix: Qualification and Validation (effective 1 December 2015)

EU GMP Annex 11: Computerised Systems (2011)

EU GMP Annex 15: Qualification and Validation (2015)

21 CFR Part 11: Electronic Records; Electronic Signatures

FDA Guidance: Part 11, Electronic Records; Electronic Signatures — Scope and Application

PIC/S PI 041-1: Good Practices for Data Management and Integrity in Regulated GMP/GDP Environments (2021)

ISPE GAMP 5, Second Edition (2022), used as non-binding industry guidance

# 1. Executive conclusion

The validation programme provides documented evidence that QCLabOne 2.0 LIMS/LES is fit for its approved GMP intended use at Suzhou Sterile Products Manufacturing Site. The system may be released to production when the go-live readiness checklist and release approval are complete.

# 2. Validation scope completed

Supplier qualification; URS; risk; architecture/design/configuration; custom development; FAT/SAT; IQ/OQ/PQ; interface, migration, access, audit trail, electronic signature, backup/restore, performance, DR and report verification; traceability; SOP/training and production baseline preparation.

# 3. Evidence summary

| Area | Conclusion |
| --- | --- |
| Requirements | 72 approved; 72 with passing final coverage |
| Tests | 98 catalogue tests executed; all pass after controlled retest |
| Defects | 6 closed; none Critical/High |
| Deviations | 4 closed; no production-data impact |
| Migration readiness | Trial criteria met; production migration plan approved |
| Part 11/Annex 11 | Applicable controls assessed and verified |
| Supplier | Qualified; audit CAPAs closed; agreement effective |
| Backup/DR | RPO 7 min and RTO 6 h 35 min demonstrated |
| Training/SOP | Prepared for completion before access/go-live |

# 4. Residual risk

Residual risks are Medium or Low. Severe consequences remain theoretically possible for calculation, CDS transfer, audit trail, backup and migration, but verified preventive/detective controls reduce likelihood and improve detection. QA and system/business owners accept the residual risk.

# 5. Open release conditions

1. Complete production migration and reconciliation. 2. Confirm production readiness checklist. 3. Complete role training and access approval. 4. Verify production baseline/checksum. 5. Obtain formal production release approval.

# 6. Validation decision

Approved for controlled production release subject to the five conditions above. No retrospective validation activity is required.

# 7. Related documents

| Relationship | Document ID | Document |
| --- | --- | --- |
| Input | QCL-2026-VP-001 | [QCLabOne 2.0 Project Validation Plan](015_Project_Validation_Plan.md) |
| Input | QCL-2026-QTA-001 | [Quality/Technical Agreement and Service Level Agreement](020_Quality_Technical_Agreement_and_SLA.md) |
| Input | QCL-2026-URS-001 | [User Requirements Specification](022_User_Requirements_Specification.md) |
| Input | QCL-2026-FRA-001 | [Functional Risk Assessment](038_Functional_Risk_Assessment.md) |
| Input | QCL-2026-TSR-001 | [Test Summary Report](063_Test_Summary_Report.md) |
| Input | QCL-2026-RTM-002 | [Final Requirements Traceability Matrix](064_Final_Requirements_Traceability_Matrix.md) |
| Output | SOP-QC-LIMS-001 | [QCLabOne 2.0 User Operation SOP](066_User_Operation_SOP.md) |
| Output | SOP-IT-LIMS-001 | [QCLabOne 2.0 System Administration SOP](067_System_Administration_SOP.md) |
| Output | SOP-QA-ATR-001 | [QCLabOne 2.0 Audit Trail Review SOP](068_Audit_Trail_Review_SOP.md) |
| Output | SOP-IT-LIMS-002 | [QCLabOne 2.0 Operations, Maintenance, Backup, Incident, Change and Security SOP](069_Operations_Maintenance_Backup_Incident_Change_and_Security_SOP.md) |
| Output | SOP-QA-ARC-001 | [Data Retention, Archiving and Record Retrieval SOP](070_Data_Retention_Archiving_and_Record_Retrieval_SOP.md) |
| Output | QCL-2026-TRN-001 | [Training Plan, Materials and Training Record](071_Training_Plan_Materials_and_Training_Record.md) |
| Output | QCL-2026-CUT-001 | [Cutover, Rollback and Contingency Plan](072_Cutover_Rollback_and_Contingency_Plan.md) |
| Output | QCL-2026-PDM-001 | [Production Data Migration Execution and Reconciliation Record](073_Production_Data_Migration_Execution_and_Reconciliation_Record.md) |
| Output | QCL-2026-PRC-001 | [Production Readiness Checklist](074_Production_Readiness_Checklist.md) |
| Output | QCL-2026-GLA-001 | [Go-Live and Production Release Approval](075_Go_Live_and_Production_Release_Approval.md) |
| Output | QCL-2026-HOV-001 | [System Handover and Support Model](076_System_Handover_and_Support_Model.md) |
| Output | QCL-2026-PBL-001 | [Production Baseline and Approved Configuration Record](077_Production_Baseline_and_Approved_Configuration_Record.md) |
| Output | QCL-2028-PRR-001 | [Annual Periodic Review Report](092_Annual_Periodic_Review_Report.md) |
| Output | QCL-2028-FDI-001 | [Final Validation and Retirement Dossier Index](108_Final_Validation_and_Retirement_Dossier_Index.md) |

# Approval record

| Approval step | Role | Case outcome |
| --- | --- | --- |
| Prepared by | System Owner | Completed |
| Reviewed by | QC Business Process Owner | Completed |
| Approved by | Site Quality Head | Approved in fictional case |

# Revision history

| Version | Date | Change |
| --- | --- | --- |
| 1.0 | 2027-06-05 | Initial approved fictional case version. |